## Coversheet
INTRODUCTION This is a list of the biochemistry, endocrinology, drugs of abuse and therapeutic drug monitoring tests performed at Eurofins Biomnis’ Dublin Laboratory. For a searchable list of tests performed by Eurofins Biomnis in France, in our laboratories in Lyon and Paris, click here https://www.eurofins-biomnis.com/en/services/test-guide/ If you cannot find details of a test you require, please contact our Client Services department on Free Phone 1800-252-966, or e-mail clientservices@ctie.eurofinseu.com All reference ranges listed are Adult Reference ranges. Paediatric reference ranges are available on request. For sample collection, please contact our Logistics department on Free Phone 1800-252-967, or e-mail llogistics@ctie.eurofinseu.com
NOTES ON SAMPLE STABILITY The majority of incorrect laboratory test results are due to improper sample collection and transport. For details regarding correct phlebotomy technique and our patient identification requirements, please click here. In order for you to arrange and properly time phlebotomy and sample collection, we have indicated, for each test, its stability after collection. Stability is indicated for whole blood at various temperatures, and for plasma or serum separated from cells, also at various temperatures. Note: RT = room temperature, i.e. 16 – 25 °C. Stability data are taken from the manufacturers’ instructions for use (IFUs), and from the World Health Organisation publication indicated below1. Sample stability data is not available for all tests under all conditions, either in the manufacturers’ IFUs or the published literature. If no information is available, in general, unless otherwise specified (such as when the required sample is whole blood), serum should be centrifuged and separated from cells after completion of clotting (20 – 30 minutes), and transported to the laboratory at 2 – 8 °C. Plasma may be centrifuged and separated from cells immediately after sampling and gently mixing the sample by inverting the tube 10 times. It should then be transported to the laboratory at 2 – 8 °C. Whole blood should be transported at 2 – 8 °C and reach the laboratory as soon as possible. However, please check each test for specific stability information. If in doubt, please contact our Client Services department on Free Phone 1800-252-966, or e-mail client.services@eurofins-biomnis.ie. References: 1. World Health Organisation: Use of anticoagulants in diagnostic laboratory investigations. WHO/DIL/LAB99.1 Rev.2, 2002. 2. Clinical Biochemistry 45 (2012) 464–469.
REASONS FOR REJECTION OF SAMPLES/NON-REPORTING OF TESTS 1. Samples received beyond the stability limits and/or not at the correct temperature indicated below for each test. 2. Samples received in the incorrect tube/with the incorrect anticoagulant or lack of the correct anticoagulant. 3. Samples received without the necessary patient identifiers. For more details, see here. 4. Samples which fail specific criteria for certain tests. See individual tests for details. 5. Leaking specimen received. 6. Samples with insufficient volume will be rejected. DRUGS OF ABUSE: For workplace Chain of Custody specimens the following criteria also apply: No seals on either specimen Seal on A container broken or tampered with Seal on B container broken or tampered with Only one specimen received Insufficient specimen for complete analysis (IA & GC-MS)

## Chemistry Assays A-Z
Analyte Name | Units | Suitable Specimen/ Container Types | Sample Stability | Turn Around Time | Instrument/ Platform/ Method | Reference Range |  | Source | Accreditation status | Note:
6MAM Heroin Metabolite | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c Cedia assay | Positivity Cut-off |  | Cedia 6MAM IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Alanine Amino Transferase ALT | U/L | Serum Serum Separator Plasma: Dipotassium EDTA Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 3 days 2-8°C: 7 days -20°C: 60 days | 24 hours | Abbott Alinity c Spectrophotometry NADH(without P-5’-P) | Adult Male | < 45 | Abbott IFU | Accredited
[6 empty] | Adult Female | < 34
Albumin | g/L | Serum: Serum separator Plasma : Dipotassium EDTA Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c Spectrophotometry (Bromocresol Green) | Adults | 35 - 50 | Abbott IFU | Accredited
[6 empty] | 60 – 90 years | 32 - 46
[6 empty] | > 90 years | 29 - 45
Globulin | g/L | Serum: Serum separator Plasma : Dipotassium EDTA Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 3 months | 24 hours | CALCULATION based on Abbott Alinity methodologies for Total Protein and Albumin. | Adult | 21 - 36 | Abbott IFU | Accredited | FORMULA: (Globulin = Total Protein - Albumin)
Alkaline Phosphatase | U/L | Serum: Serum separator Plasma: Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 7 days | 24 hours | Abbott Alinity c Spectrophotometry Hydrolysis of Para- Nitrophenyl phosphate | 16 to 21 years (Male) | 56 - 167 | Abbott IFU | Accredited
[6 empty] | 16 to 29 years (Female) | 44 - 107
[6 empty] | 22 to 79 years (Male) | 50 - 116
[6 empty] | 30 to 79 years (Female) | 46 - 122
Alpha-1 Antitrypsin A1AT | g/L | Serum : Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Sodium heparin Potassium EDTA Sodium citrate | 2-8°C: 2 days -20°C: not specified | 24 hours | Abbott Alinity c Immunoturbidimetry |  | 0.9 - 2.0 | Abbott IFU | Accredited
Alpha- FetoProtein AFP | IU/mL | Serum: Serum, Serum separator. Plasma: Sodium heparin, Lithium heparin, Dipotassium EDTA Sodium EDTA | 20-25°C: 3 days 2-8°C: 7 days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 0.74 - 7.29 | Abbott IFU | Accredited
Amikacin | mg/L | Serum: Serum tubes (with or without gel barrier) Plasma- Acceptable anticoagulants are: Lithium heparin, Sodium heparin K2-EDTA, K3-EDTA | 2-8°C: 7 days -20°C: 14 days | 24 hours | Abbott Alinity c Homogeneous particle-enhanced turbidimetric inhibition immunoassay (PETINIA). | Trough | 4 - 8 | Abbott IFU | Accredited
[6 empty] | Severe Infection (Peak) | 25 - 35
[6 empty] | Toxic Levels (Peak) | >35
Amphetamine/ Meth amphetamine - Urine | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c Enzymatic Immunoassay | Positivity Cut-off | 1000 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Amylase 2 | U/L | Serum: Serum tubes (with or without gel barrier) Plasma- Acceptable anticoagulants are: Lithium heparin Sodium heparin | 20-25°C: 24 hr 2-8°C: 7 days -20°C: 3 month | 24 hours | Abbott Alinity c Enzymatic + Colorimetric | Pediatric |  | Abbott IFU | Accredited
[6 empty] | 0-14 days | 3 - 10
[6 empty] | 15 day to < 13 week | 2 - 22
[6 empty] | 13 week to < 1 year | 3 - 50
[6 empty] | 1 year to < 18 year | 25 - 101
[6 empty] | Adult | 28 - 100
Angiotensin Converting Enzyme (ACE) | U/L | Serum is the preferred specimen. Plasma: Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c Enzymatic/ Colorimetric Hydrolysis of Furylacryloylphenylalanylglyclglycine (FAPGG) | >14 years | 8 - 65 | Glenbio IFU | Accredited | Serum, which has been separated from the cells as soon as possible after collection, is the only suitable sample type. ACE is a zinc-dependant enzyme and anticoagulants, especially EDTA, can lead to falsely low results.
Anti-TG | IU/mL | Serum: Serum, Serum separator Plasma: Lithium heparin Potassium EDTA | 20-25°C: 8hours 2-8°C: 72 hours -20°C: 30days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | < 4.11 | Abbott IFU | Accredited
Anti-TPO | IU/mL | Serum: Serum, Serum separator Plasma: Lithium heparin Potassium EDTA | 20-25°C: 8hours 2-8°C: 72 hours -20°C: 30days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | < 5.61 | Abbott IFU | Accredited
Aspartate Amino transferase AST | U/L | Serum Serum Separator Plasma: Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 4 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c Spectrophotometry | Adults | 11- 34 | Abbott IFU | Accredited
Barbiturates, semiquantitative | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c Enzymatic Immunoassay | Positivity Cut-off | 200 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Benzodiazepines semiquantitative | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c Enzymatic Immunoassay | Positivity Cut-off | 200 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Beta-2-Microglobulin | mg/L | Plasma Collection tubes Acceptable anticoagulants are: EDTA Sodium heparin | 2-8°C: 8 days | 24 hours | Abbott Alinity c Turbidimetric/ Immunoturbidimetry | Adult | 0.97 - 2.64 | Abbott IFU | Accredited
Βeta-HCG, Total | mIU/mL | Serum: Serum, Serum separator Plasma: Dipotassium EDTA, Tripotassium EDTA Lithium heparin Lithium heparin plasma separator Sodium heparin | 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Males and Non- Pregnant Females | < 5.0 | Abbott IFU | Accredited | In patients receiving therapy with high biotin doses no sample should be taken until at least 8 hours after the last biotin administration.
Pregnancy- Weeks LMP
[6 empty] | 1-10 | 202 - 231000
[6 empty] | 11- 15 | 22536 - 234990
[6 empty] | 16 - 22 | 8007 - 50064
[6 empty] | 23 - 40 | 1 600 - 49 413
Bile acid | μmol/L | Serum Plastic tubes (with or without gel barrier) Plasma- Acceptable anticoagulants are: Lithium Heparin (with or without gel barrier) Sodium Heparin K2EDTA | 20-25°C: 1 day 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c Cyclic-enzymatic/ colorimetric | Adult | 1.0 - 6.0 | Abbott IFU | Accredited | The blood sample must be collected before the administration of Ursodeoxycholic Acid. ideally, sample should be taken after an 8-hour fast.
[6 empty] | Pregnancy | 1.3 - 9.0
Bilirubin 2 Total | µmol/L | Serum: Serum separator Plasma: Lithium heparin and Lithium heparin separator Sodium heparin Dipotassium EDTA | 20-25°C: 8 hr 20-25°C: 24 hr (serum separator and Lithium heparin separator) 2-8°C: 7 days -20°C: 3 mth | 24 hours | Abbott Alinity c Spectrophotometry, Diazonium salt | Premature: | 0 - 1day <136.8 1-2 day <205.2 3-5 day <273.6 | Abbott IFU (revised Dec 2023) | Accredited | Specimens should be protected from bright light as bilirubin is photolabile. Tubes containing sodium fluoride/ potassium oxalate is not recommended due to the potential for hemolysis with this anticoagulant.
[6 empty] | Full-term | 0 - 1day 34.2 - 102.6 1-2 day 102.6 - 171 3-5 day 25.7 - 205.2
[6 empty] | Premature: Full-term Adult | 5.1 - 20.5
Bilirubin Direct | µmol/L | Serum: Serum tubes (with or without gel barrier) Plasma - Acceptable anticoagulants are:Lithium heparin Lithium heparin (with or without gel barrier) Sodium heparin EDTA | 20-25°C: 2 days 2-8°C: 7 days -20°C: 3months | 24 hours | Abbott Alinity c Spectrophotometry, Diazo-Reaction | Adult | 0.0 - 8.6 | Abbott IFU | Accredited
Bilirubin, Indirect |  | Serum: Serum tubes (with or without gel barrier) Plasma- Acceptable anticoagulants are: Lithium heparin, Sodium heparin, EDTA | 20-25°C: 8hours 2-8°C: 7 days -20°C: 3months | 24 hours | CALCULATION based on Abbott Alinity methodologies for Total Bilirubin and Direct Bilirubin | Adult | 3.0 - 12.0 | Abbott IFU | Accredited | FORMULA: (Indirect Bilirubin = Total Bilirubin - Direct Bilirubin)
CA 125 II | U/mL | Serum: Serum, Serum separator. Plasma: Tripotassium EDTA, Sodium Heparin, Lithium Heparin | 2-8°C: 7days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | < 35 | Abbott IFU | Accredited
CA 15-3 | U/mL | Serum: Serum, Serum separator. Plasma: Tripotassium EDTA, Sodium Heparin, Lithium Heparin | 2-8°C: 7days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | < 31.3 | Abbott IFU | Accredited
CA 19-9XR | U/mL | Serum: Serum, Serum separator. Plasma: Tripotassium EDTA, Sodium Heparin, Lithium Heparin | 2-8°C: 7days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | <37 | Abbott IFU | Accredited
Calcium 2 | mmol/L | Serum : Serum tubes Serum separator tubes . Plasma: Lithium heparin, Sodium heparin . Urine: Random or 24 hour | Serum: 20-25°C: 3 days 2-8°C: 3 weeks -20°C: 8months . Urine: 20-25°C: 4 days 2-8°C: 7 days -20°C: 3months | 24 hours | Abbott Alinity c Spectrophotometry, Arsenazo III | Serum: |  | In-house study "September 2022, PSM version 5.02" as per Dr Michael Louw, 26-08-2025 | Accredited | Urine: use clean urine collection container. Keep specimens on ice during collection Random urine: collect specimen in a bottle containing 1 to 2 mL of 6 mol/L HCL to prevent calcium salt precipitation. 24 hour Urine: collect specimen in a bottle containg 20 to 30 mL of 6 mol/L HCL to prevent calcium salt precipitation.
[6 empty] | upto 1 month | 1.9 - 2.5
[6 empty] | upto 1 year | 2.1 - 2.7
[6 empty] | upto 4 year | 2.1 - 2.6
[6 empty] | upto 20 year | 2.3 - 2.75
[6 empty] | Adult | 2.1 - 2.6
[6 empty] | Urine: |  | Abbott IFU
[6 empty] | Pediatric | Up to 0.15 mmol/kg/day
[6 empty] | Adult Ca-free diet | 0.13 - 1.00 mmol/day
[6 empty] | Adult Ca diet | 2.50 - 7.50 mmol/day
[6 empty] | Random - male | 0.23 - 9.48 mmol/L
[6 empty] | Random - female | 0.13 - 8.93 mmol/L
Calcium, Calculated | mmol/L |  |  | 24 hours | CALCULATION based on Abbott Alinity methodologies for Calcium and Albumin. |  |  | Abbott IFU | Accredited | FORMULA: Calcium Corrected = ((40 - Albumin)*0.02) + Calcium))
Cannabinoids, semiquantitative | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c Enzymatic Immunoassay | Positivity Cut-off | 50 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Carbon Dioxide/ Bicarbonate | mmol/L | Serum : Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Lithium heparin Sodium heparin | 20-25°C:2hours 2-8°C: 2 days -20°C: 2 weeks | 24 hours | Abbott Alinity c Spectrophotometry, PEP Carboxylase | Adult | 22 to 29 | Abbott IFU | Accredited | Note: Bicarbonate content in uncapped tubes decreases approximately 4 mmol/L after one hour. Serum stored in open tubes is stable for up to 4 hours. Samples must be run immediately once uncapped.
[6 empty] | > 60 years | 23 to 31
Carcino Embryonic Antigen CEA | ng/mL | Serum: Serum, Serum separator. Plasma: Sodium heparin, Lithium heparin, Potassium EDTA | 2-8°C: 7days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | <5 | Abbott IFU | Accredited
Ceruloplasmin | g/L | Serum: Serum tubes (not gel tubes) Plasma: Lithium heparin (not gel tubes) | 20-25°C: 8 days 2-8°C: 2 weeks -20°C: 3months | 24 hours | Abbott Alinity c Turbidimetric/ Immunturbidimetric | Adult | 0.2 to 0.6 g/L | Abbott IFU | Accredited
Chloride | mmol/L | Serum : Serum tubes (with or without gel barrier) For Potassium, hemolyzed specimens must not be used. Plasma -Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c Indirect ISE | Adult | 98 - 107 | Abbott IFU | Accredited | Must be seperated ASAP
Chloride-Urine | mmol/ 24hr | Urine (timed, 24- hour) Without preservatives | 20-25°C: 7 days 2-8°C: 7 days -20°C: 7 days | 24 hours | Abbott Alinity c Indirect ISE | Adult <60 years | 110 - 250 | Abbott IFU | Accredited
[6 empty] | Adult >60 years | 95 - 195
| mmol/L | Spot urine (random) |  | 24 hours |  | No reference range available
Cholesterol, Total | mmol/L | Serum: Serum separator Plasma Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c Enzymatic | Adult | < 5.0 | Source: ESC/EAS Guidelines for the management of dyslipidaemias. http://www.eas-society.org/guidelines-2.aspx; Abbott IFU | Accredited | Fasting sample required
Cholesterol, LDL (Direct) | mmol/L | Serum : Serum tubes (with or without gel barrier) Plasma - Acceptable anticoagulants are: Lithium heparin, Sodium heparin, EDTA Anticoagulants containing citrate should not be used. | 2-8°C: 5 days -80°C: 3 months | 24 hours | Abbott Alinity c, Selective resolution of LDL-Particles under dye formation |  | 0.0 – 3.0 | Source: ESC/EAS Guidelines for the management of dyslipidaemias. http://www.eas-society.org/guidelines-2.aspx; Abbott IFU | Accredited | Fasting sample required. Separate plasma from red blood cells or gel as soon after collection as possible (within 3hours)
Cholesterol, Ultra HDL | mmol/L | Serum : Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin, Sodium heparin, EDTA | 20-25°C: 2 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c, Accelerated enzymatic Reaction/selective solvent | Adult | >1.0 | National Cholesterol Education Program (NCEP) Adult Treatment Panel III Report and European Guidelines http://www.eas-society.org/guidelines-2.aspx; Abbott IFU | Accredited | Fasting sample required.
Cocaine, semiquantitative | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c, Enzyme immunoassay | Positivity Cut-off | 300 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Complement C3 | g/L | Serum: Serum tubes (with or without gel barrier) Plasma - Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin | 2-8°C: 3 days -20°C: 8 days | 24 hours | Abbott Alinity c Immunoturbidimetry | Adult Male 14 - 80 years | 0.82 - 1.85 | Abbott IFU | Accredited
[6 empty] | Adult Female 14 - 80 years | 0.83 - 1.93
Complement C4 | g/L | Serum : Serum tubes (with or without gel barrier) Plasma - Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin EDTA | 2-8°C: 2 days -20°C: 2 days | 24 hours | Abbott Alinity c Immunoturbidimetry | Adult Male 14 - 80 years | 0.15 - 0.53 | Abbott IFU | Accredited
[6 empty] | Adult Female 14 - 80 years | 0.15 - 0.57
Cortisol | nmol/L | Serum: Serum, Serum separator. Plasma: Lithium Heparin, Sodium Heparin, Potassium EDTA Plasma separator tubes with Lithium Heparin. | 2-8°C: 14 days -20°C: 30days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Morning 08H00 | 171 - 800 | Handbook of Diagnostic Endocrinology, 2nd Edition, 2008. William E Winter & al. AACC Press; Abbott IFU | Accredited | Time of collection should be written on the tube as the time of injection in the case of dynamic test (Synacthen / Dexamethasone test). Allow serum to clot completely at room temperature. Separate serum or plasma from cells ASAP.
[6 empty] | Afternoon | Approx. half the morning values
In the evaluation of Adrenal failure
[6 empty] | Highly Unlikely | >550
[6 empty] | Virtually Diagnostic | <138
Evening 24H00: In the evaluation of Cushing's Syndrome
[6 empty] | Virtually excludes | <138
[6 empty] | Highly suggestive | <207
Cortisol - Urinary | nmol/24hr | Urine: The urine sample must be collected in a clean, previously unused container. Preservatives are not required; however, ten grams of boric acid per liter of urine may be used. | 2-8°C: 14 days -20°C: 30days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 12 - 486 | Abbott IFU | Accredited
C-peptide | ug/L | Serum: Serum, Serum separator. Plasma: Potassium EDTA, Lithium heparin, Sodium heparin, Ammonium heparin Sodium fluoride / potassium oxalate Plasma separator (lithium heparin) | 20-25°C: 24hrs 2-8°C: 48hrs -20°C: 3 months | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult (fasting, 12hours) | 0.80-5.20 | Abbott IFU | Accredited | Overnight fast.
[6 empty] | Adult (post-prandial) | 2.0 - 9.0
C-Reactive Protein - High Sensitivity | mg/L | Serum: Serum tubes (with or without gel barrier) Plasma - Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin EDTA | 20-25°C: 15days 2-8°C: 2 months -20°C: 1 year | 24 hours | Abbott Alinity c, Turbidimetric/ Immunturbidimetric | Adult | < 5.0 | Pearson TA et al, Circulation 2003; 107:499-511; Abbott IFU | Accredited
[6 empty] | Lowest relative CVD risk | < 1.0
[6 empty] | Average relative CVD risk | 1.0 - 3.0
[6 empty] | Highest relative CVD risk | > 3.0
Creatine kinase CK | U/L | Serum Serum tubes (with or without gel barrier) Plasma - Acceptable anticoagulants: Lithium heparin Sodium heparin | 20-25°C: 2 days 2-8°C: 7 days | 24 hours | Abbott Alinity c, NAC (N-Acetyl-L-Cystein) | Adult Male | 30 - 200 | Abbott IFU | Accredited
[6 empty] | Adult Female | 29 - 168
Creatinine Enzymatic | µmol/L | Serum Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin EDTA is not recommended. | 20-25°C: 7 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c, Enzymatic | Adult Male | 64 - 104 | Abbott IFU | Accredited
[6 empty] | Adult Female | 49 - 90
Creatinine Enzymatic -Urine | mmol/ 24hr | Urine -Timed specimens collected over intervals shorter than 24 hours | 20-25°C: 2 days 2-8°C: 6 days -20°C: 6 months | 24 hours | Abbott Alinity c, Enzymatic | Adult Male | 7.7 - 21.3 | Abbott IFU | Accredited
[6 empty] | Adult Female | 5.9 - 14.1
|  | Urine (random/spotspecimens) Clean plastic or glass container without preservatives |  |  |  |  | N/A
N/A
EGFR |  | Serum Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin EDTA is not recommended. | 20-25°C: 7 days 2-8°C: 7 days -20°C: 3 months | 24 hours | CALCULATION based on Creatinine methodology and patient demographics. | Adult Male | > 90: Normal 60 - 89: Normal or Stage 2 CKD 30 - 59: Moderate Impairment, Stage 3 CKD 15 - 29: Severe Impairment, Stage 4 CKD <15: Established renal failure, Stage 5 CKD | https://kdigo.org/ | Accredited | FORMULA: EGFR= ((Creatinine^1.154)*32788)*1) / (Age ^0.203)
[6 empty] | Adult Female |  |  | Accredited | FORMULA: EGFR= ((Creatinine^1.154)*32788)* 0.742)*1)) / (Age ^0.203)
DHEA-S | µmol/L | Serum: Serum, Serum separator. Plasma: Potassium EDTA, Sodium citrate Sodium heparin | 2-8°C: 8days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Male: 19Y | 1.2 -10.4 | Abbott IFU | Accredited
[6 empty] | Male: 24Y | 6.5 -14.6
[6 empty] | Male: 34Y | 4.6 -16.1
[6 empty] | Male: 44Y | 3.8 -13.1
[6 empty] | Male: 54Y | 3.7 -12.1
[6 empty] | Male: 64Y | 1.3 -9.8
[6 empty] | Male: >64 | 6.2 -7.7
[6 empty] | Female: 19Y | 1.7 -13.4
[6 empty] | Female: 24Y | 3.6 -11.1
[6 empty] | Female: 34Y | 2.6 -13.9
[6 empty] | Female: 44Y | 2 -11.1
[6 empty] | Female: 54Y | 1.5 -7.7
[6 empty] | Female: 64Y | 0.8 -4.9
[6 empty] | Female: >64 | 0.9 -2.1
Digoxin | ug/L | Serum : Serum tubes (with or without gel barrier) Plasma - Acceptable anticoagulants are: Lithium heparin Sodium heparin Potassium EDTA Heparin gel plasma separator | 2-8°C: 48 hours -20°C: 7 days | 24 hours | Abbott Alinity c, Particle-enhanced turbidimetric inhibition immunoassay (PETINIA) |  | 0.6 – 1.2 >2 associated with toxicity | Therapeutic range for digoxin as recommended by European Society of Cardiology Guidelines for the diagnosis and treatment of acute and chronic heart failure. European Heart Journal (2008) 29, 2388 - 2442.); Abbott IFU | Accredited
Ecstasy, semiquantitative | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c, Enzyme immunoassay | Positivity Cut-off | 500 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
EDDP Methadone Metabolite | g/ml | Urine Clean plastic or glass container | 20-25°C: 7 days 2-8°C: 2months -20°C: >2months | 24 hours | Abbott Alinity c Immunalysis assay | Positivity Cut-off | 100 | Immunalysis EDDP IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Estradiol | pmol/L | Serum: Serum, Serum separator Plasma: Lithium heparin, Plasma separator, Potassium EDTA | 2-8°C: 7 days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Follicular phase | 77 – 922 | Abbott IFU | Accredited | In patients receiving therapy with high biotin doses no sample should be taken until at least 8 hours after the last biotin administration.
[6 empty] | Mid Cycle phase | 140 – 2383
[6 empty] | Luteal phase | 77 – 1145
[6 empty] | Post- meno pausal no HRT | < 103
[6 empty] | Post- meno pausal on HRT | < 529
[6 empty] | Male | 40 – 161
Ethanol | mg/dL | Urine Clean plastic or glass container | 2-8°C: 30 days -20°C:longer | 24 hours | Abbott Alinity c Enzymatic (Alcohol Dehydrogenase) | Positivity Cut-off | 10 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Faecal Immunochemical Test (FIT) | ng/ml | Faeces sample collected into an analyser specific sampling bottle containing buffer (OC Auto Sampling Bottle 3). | 2-8°C: 10 days | 24 hours | OC-Sensor Pledia, Latex Agglutination Immunoturbidimetry | Normal range cut-off | <100 | BowelScreen Guidelines for Quality Assurance in Colorectal Screening, Second Edition, Published 2017, ISBN 978-1-907487-26-2 | Accredited
[6 empty] | NSS Bowel Screen Normal Range Cut-off | <225
Ferritin | ng/mL | Serum: Serum, Serum separator Plasma: Tripotassium EDTA, Lithium heparin | 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult Male | 15 - 200 | Abbott IFU | Accredited | Reference ranges: According to WHO guidelines and CALIPER database.
[6 empty] | Adult Female | 15 - 150
[6 empty] | 4 – 14 days M/F | 12 – 717
[6 empty] | 15 days to < 6 months M/F | 12 - 647
[6 empty] | 6 months to < 1 year M/F | 12 – 182
[6 empty] | 1 to < 5 years M/F | 12 – 100
[6 empty] | 5 to < 14 years M/F | 15 – 79
[6 empty] | 14 to < 19 years F | 15 - 67
[6 empty] | 14 to < 16 years M | 15 - 83
[6 empty] | 16 to < 19 years M | 15 – 172
Folate | ng/mL | Folate: Serum: Serum, Serum separator Plasma: Lithium heparin plasma, Lithium heparin plasma separator. | 2-8°C: 7 days -20°C: 30 days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 3.1 - 20.5 | Abbott IFU | Accredited | Protect from light
Free T3 | pmol/L | Serum: Serum, Serum separator Plasma: Sodium heparin, Lithium heparin, Potassium EDTA | 2-8°C: 6 days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 2.42 - 6.01 | Abbott IFU | Accredited
Free T4 | pmol/L | Serum: Serum, Serum separator Plasma: Sodium heparin, Lithium heparin, Lithium heparin plasma separator, Potassium EDTA | 2-8°C: 6 days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 9.01 - 19.05 | Abbott IFU | Accredited | In patients receiving therapy with high biotin doses no sample should be taken until at least 8 hours after the last biotin administration.
Follicle Stimulating Hormone FSH | UI/L | Serum: Serum, Serum separator Plasma: Sodium heparin, Lithium heparin, Potassium EDTA | 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Follicular phase | 3.0 - 8.1 | Abbott IFU | Accredited | In patients receiving therapy with high biotin doses no sample should be taken until at least 8 hours after the last biotin administration.
[6 empty] | Mid Cycle phase | 2.6 - 16.7
[6 empty] | Luteal phase | 1.4 - 5.5
[6 empty] | Post-Meno pausal | 26.7 - 133
[6 empty] | Males | 1.0 - 12.0
Gamma-Glutamyl Transferase | U/L | Serum Serum separator Plasma: Lithium heparin, Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c, L-Gammaglutamyl-3-carboxy-4-nitroanilid-Substrate | Male | < 55 | Abbott IFU | Accredited
[6 empty] | Female | < 38
Gentamicin | mg/L | Serum : Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin K2-EDTA K3-EDTA | 2-8°C: 7 days <-10°C: 14 days | 24 hours | Abbott Alinity c, Particle-enhanced turbidimetric inhibition immunoassay (PETINIA) | Trough (Less Severe Infection) | <1 | Abbott IFU | Accredited | Please include dose regime (dose and frequency of administration, and timing of sampling (trough and/or peak)).
[6 empty] | Trough (Severe Infection) | <2 - 4
[6 empty] | Peak (Less Severe Infection) | 5 - 8
[6 empty] | Peak (Severe Infection) | 8 - 10
[6 empty] | Toxic Levels | >10 - 12
Glucose | mmol/L | Plasma -Acceptable anticoagulants are: Sodium Fluoride Potassium Oxalate Sodium Fluoride/K2 EDTA | 20-25°C: 2 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c, Enzymatic (Hexokinase/ G-6-PDH) | Fasting glucose | ≤ 6.00 | WHO criteria for the diagnosis of diabetes mellitus. Abbott IFU | Accredited | Fasting sample required.
Haemoglobin A1C (IFCC) | mmol/mol | Blood samples collected in primary tubes containing K2EDTA may be stored at 25 °C for 24 hours or at 4 °C for 14 days before analysis. | 20-25°C: 1 day 4°C: 14 days | 24 hours | TOSOH High Performance Liquid Chromatography (HPLC) | Adult | 20 - 42 | Irish College of General Practitioners "A Practical Guide to Integrated Type II Diabetes Care 2016" | Accredited
Homocysteine | µmol/L | Serum: Serum, Serum separator Plasma: Lithium heparin, Potassium EDTA | ON ICE: 6hours 2-8°C: 14 days -20°C:1 year | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult Male | 5.5 - 16.2 | Abbott IFU | Accredited | To minimize increases in homocysteine concentration from synthesis by red blood cells, place all specimens (serum and plasma) on ice after collection and prior to processing.
[6 empty] | Adult Female | 4.4 - 13.6
Immunoglobulin IgA | g/L | Serum : Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin EDTA | 20-25°C: 7 days 2-8°C: 7 days -20°C: 6 months | 24 hours | Abbott Alinity c, Immuno turbidimetric | Male 12 to 60 years | 0.63 - 4.84 | Abbott IFU | Accredited
[6 empty] | Female 12 to 60 years | 0.65 to 4.21
[6 empty] | Male > 60 years | 1.01 to 6.45
[6 empty] | Female > 60 years | 0.69 to 5.17
Immunoglobulin IgE | IU/mL | Serum: Serum tubes Plasma -Acceptable anticoagulants are: Sodium heparin, Lithium heparin, Sodium EDTA Potassium EDTA, Sodium citrate | 2-8°C: 2 days | 24 hours | Abbott Alinity c, Immuno turbidimetric | Adults | < 100 IU/mL | Abbott IFU | Accredited
Immunoglobulin IgG | g/L | Serum : Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 6 months | 24 hours | Abbott Alinity c, Immuno turbidimetric | Male 2 to 80 years | 5.4 - 18.22 | Abbott IFU | Accredited
[6 empty] | Female 2 to 80 years | 5.52 - 16.31
Immunoglobulin IgM | g/L | Serum : Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin EDTA | 20-25°C: 7 days 2-8°C: 7 days -20°C: 6 months | 24 hours | Abbott Alinity c, Immuno turbidimetric | Male >12years | 0.22 - 2.40 | Abbott IFU | Accredited
[6 empty] | Female >12years | 0.33 - 2.93
Insulin | mIU/L | Serum: Serum, Serum separator Plasma: Potassium EDTA Sodium EDTA, Sodium heparin, Sodium fluoride | -20°C: 7 days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Fasting | 3.4 - 19.6 | Williams Textbook of Endocrinology 13th edition 2015; Abbott IFU | Accredited | Overnight fast. Provide fresh samples if possible. NB: please note that insulin is unstable in whole blood. Serum or plasma must be separated from red cells within 30 minutes. minutes of sampling. Failure to do so may lead to falsely low results.
[6 empty] | Post-prandial | 3.0 - 50.0
Intact Parathyroid Hormone PTH | pg/mL | Serum: Serum ( use of serum separator tubes may result in a decrease in concentration ) Plasma: Potassium EDTA, Lithium Heparin, Sodium Heparin | 2-8°C: 2 days -20°C: 6 months | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 15 - 68.3 | Abbott IFU | Accredited | Fasting sample required.
Iron | µmol/L | Serum : Serum separator Plasma -Acceptable anticoagulants are: Lithium Heparin Sodium heparin | 20-25°C: 10hours 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c, Ferene | Adult (Female) | 9.0 - 30.4 | Abbott IFU | Accredited | EDTA, oxalate, or citrate as anticoagulants must not be used, since they bind iron ions, preventing its reaction with the chromogen. Specimens should be collected in the morning to avoid low results due to diurnal variation.
[6 empty] | Adult (Male) | 11.6 – 31.3
Lactate Dehydrogenase LDH | U/L | Serum: Serum separator . Plasma: Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 3days 2-8°C: 3 days -20°C: 8 weeks | 24 hours | Abbott Alinity c, Oxidation of Lactate to Pyruvate | Adult | 125 - 220 | Abbott IFU | Accredited | Erythrocyte LDH activity is 150 times that of plasma and LDH is extremely sensitive to even minor haemolysis induced by sample transport including transport by pneumatic tube systems. Serum or plasma should therefore be separated from red cells immediately after collection (plasma), or immediately after clotting of serum. Samples should ideally not be sent to the laboratory unseparated.
[6 empty] | 0 to <15 days | 309 - 1222 | https://caliper.research.sickkids.ca/#/ | Accredited
[6 empty] | 15 days to < 1 yr | 163 - 452
[6 empty] | 1 to < 10 years | 192 - 321
[6 empty] | F: 10 to 15 years | 157 - 272
[6 empty] | M: 10 to 15 years | 170 - 283
[6 empty] | 15 to 19 years | 130 - 250
Lipase | U/L | Serum: Serum separator . Plasma: Lithium heparin Sodium heparin EDTA unsuitable | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c, Kinetic Colorimetric method | Adult | 0-60 | Sentinel Diagnostics IFU | Accredited
Lipoprotein [a] | g/L | Serum: Serum tubes Plasma- Acceptable anticoagulants: EDTA | 2-8°C: 8 days | 24 hours | Abbott Alinity c, Turbidimetric/ Immunturbidimetric | Adult | <500 Values above 0.500 g/L are associated with an increased risk of athero- sclerosis. | Abbott IFU, European Atherosclerosis Society Consensus Position Paper (Lipoprotein(a) as a cardiovascular risk factor: current status. Eur Heart J (2010) doi: 10.1093/eurheartj/ehq386) | Accredited
Lithium | mmol/L | Serum: Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Sodium heparin K2-EDTA Do not use lithium heparin. | 20-25°C: 1 day 2-8°C: 7 days -20°C: 6 months | 24 hours | Abbott Alinity c, Colorimetric Method | Immediate-release formulations | 0.5 - 0.8 12hours after last dose | Abbott IFU, https://www.nice.org.uk/guidance/cg185/resources/bipolar-disorder-assessment-and-management-35109814379461) | Accredited | Sample to be taken at trough. Lithium has a narrow therapeutic range and toxicity should also be suspected even when lithium is within the target range if symptoms are present and in compromised patientse.g. older patients, interacting drugs such as NSAIDs/diuretics, sodium depletion, decreased renal function, <50 kg body weight.
[6 empty] | Slow-release formulations | 0.8 - 1.20 12hours after last dose
0.5 - 0.8 immediately before next dose
Luteinizing Hormone LH | IU/L | Serum: Serum, Serum separator Plasma: Potassium EDTA, Sodium heparin | 2-8°C: 7 days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Follicular phase | 1.8 - 11.8 | Abbott IFU | Accredited | In patients receiving therapy with high biotin doses no sample should be taken until at least 8 hours after the last biotin administration.
[6 empty] | Mid Cycle phase | 7.6 – 89.1
[6 empty] | Luteal phase | 0.6 - 14.0
[6 empty] | Post Meno pause | 5.2 – 62
[6 empty] | Male | 0.6 - 12.1
Magnesium | mmol/L | Serum : Serum tubes (with or without gel barrier) Use nonhemolyzed specimens. Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin, Sodium heparin | 20-25°C: 1 day 2-8°C: 3 days -20°C: 3 month | 24 hours | Abbott Alinity c, Enzymatic | Adult | 0.66 - 1.07 | Abbott IFU | Accredited | Plasma samples collected with EDTA anticoagulant or specimens from patients receiving EDTA are unsuitable for analysis, because this compound chelates magnesium, making it unavailable for reaction with the reagent. Sodium fluoride and oxalate also interfere with the results and should be avoided.
Magnesium-Urine | mmol/ 24hr | Urine (24 hour) Collect specimens in a container with boric acid or 20 to 30 mL of 6N HCl to prevent precipitation of magnesium complexes. | 20-25°C: 2 days 2-8°C: 2 days -22°C: 1 year | 24 hours | Abbott Alinity c, Enzymatic | Adult | 3.00 – 5.00 | Abbott IFU | Accredited | Do not use more than 2.5 mL 6N HCl per 100 mL of urine. Excess hydrochloric acid may cause elevated results with this methodology. Do not exceed 10 g/L boric acid.
| mmol/L | Spot urine | 20-25°C: 2 days 2-8°C: 2 days -22°C: 1 year | 24 hours | Abbott Alinity c, Enzymatic | No reference range
Methadone, semiquantitative | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c, Enzyme immunoassay | Positivity Cut-off | 300 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Microalbumin - Urine | mg/mmol | Urine spot: Clean, unused plastic or glass container with preservatives | 2-8°C: 6 days | 24 hours | Abbott Alinity c, Turbidimetric/ Immunturbidimetric | ACR (Albumin/ Creatinine Ratio): | <3.0 | Abbott IFU, NICE Guideline NG203, August 2021 | Accredited
|  | Urine timed/24hr: Clean, unused plastic or glass container with preservatives | 2-8°C: 3 days -70°C: 5 months | 24 hours
Albumin/ Creatinine Ratio (ACR) | mg/mmol | Urine spot/ timed/24hr: Clean, unused plastic or glass container with preservatives |  | 24 hours | CALCULATION based on Abbott Alinity methodologies for Microalbumin and Urinary creatinine. | ACR (Albumin/ Creatinine Ratio): | <3.0 | NICE Guideline NG203, August 2021 | Accredited | FORMULA: ACR = Microalbumin / Urinary creatinine
NT-pro BNP | pg/ml | Serum: Serum, Serum separator Plasma: Potassium EDTA, Lithium heparin | 20-25°C: 3 days 2-8°C: 6 days -20°C: 30days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adults <75 years | < 125.0 | Abbott IFU | Accredited
[6 empty] | Adults >75 years | < 450.0
Opiates, semiquantitative | ng/mL | Urine Clean plastic or glass container | 2-8°C: 5 days -20°C:longer | 24 hours | Abbott Alinity c, Enzyme immunoassay | Positivity Cut-off | 300 | Abbott IFU | Accredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request.
Osmolality | mOsm/kg | Serum: Serum, Serum separator Plasma: Potassium EDTA, Lithium heparin | 20-25°C: 2 days 2-8°C: 8 days -20°C: 30 days | 24 hours | Genotech Osmometer, Freezing-point depression osmometry | Adult | 280 - 298 | Clinical Chemistry 44: 1582, 1998 | Unaccredited | Whole blood stability was higher in the presence of anticoagulant. Plasma was less stable when refrigerated. Source for sample stability: https://pubmed.ncbi.nlm.nih.gov/28372954/
|  | Urine Clean plastic or glass container | 20-25°C: 5 days 4°C: 4 days -20°C: 30days |  |  | Adult | 50 - 1200 | Wu, A.H.B. ed: Tietz Clinical Guide to Laboratory Tests 4th Edition, Saunders 2006 |  | Source for sample stability: https://pubmed.ncbi.nlm.nih.gov/28372954/
Phenytoin | µg/mL | Serum Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin, Sodium heparin,Potassium EDTA, Sodium citrate, Sodium fluoride/potassium oxalate | 20-25°C: 2 days 2-8°C: 1 month -20°C: 5 month | 24 hours | Abbott Alinity c, Enzyme immunoassay | Adult | 10.0 - 20.0 | Abbott IFU | Accredited
[6 empty] | Toxicity | > 20.0
Phosphate | mmol/L | Serum: Serum tubes Serum separator tubes Plasma-Acceptable anticoagulants: Dipotassium EDTA Lithium heparin Sodium heparin | 20-25°C: 1 day 2-8°C: 3 days -20°C: 30 days | 24 hours | Abbott Alinity c, Spectrophotometry, Phosphomolybdate | Adult | 0.81 - 1.45 | Abbott IFU | Accredited | Serum and plasma should be free of fibrin, red blood cells, platelets and with any visible hemolysis. The specimen should be separated from the clot as soon as possible to prevent falsely elevated phosphate levels due to passage of phosphate from the erythrocytes into the serum. The only acceptable anticoagulant is heparin. Sample must be separated (centrifuged) within 6 hours after collection. Whole blood samples which are not separated within 6 hours will be rejected.
Phosphate- Urine | mmol/ 24hr | Urine (24 hour) Clean plastic or glass container (see special note). | 20-25°C: 4 days 2-8°C: 7 days -20°C: 1 month | 24 hours | Abbott Alinity c, Spectrophotometry, Phosphomolybdate | Adult | 12.9 - 42.0 | Abbott IFU | Accredited | 24hr Urine specimens should be collected in 6 mol/L HCl, 20 to 30 mL, to avoid precipitation of phosphate complexes.
| mmol/ L | Spot Urine (random) Clean plastic or glass container without preservatives |  |  |  | Adult: Male Female | 1.6 - 61 2.3 - 48
Potassium | mmol/L | Serum : Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin (full draw) | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c, Indirect ISE | Adult | 3.5 - 5.1 | Abbott IFU | Accredited | Must be separated ASAP. Unseparated samples >2hours old are unsuitable for analysis. Haemolyzed specimens must not be used.
Potassium-Urine | mmol/ 24hr | Urine (random; 24- hour) Without preservatives | 20-25°C: 45 day 2-8°C: 2 months -20°C: 1 year | 24 hours | Abbott Alinity c, Indirect ISE | Adult | 25 - 125 | Abbott IFU | Accredited
Progesterone | nmol/L | Serum: Serum, Serum separator Plasma: Sodium heparin, Lithium heparin Potassium EDTA | 2-8°C: 10 days -20°C: 6 months | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Follicular Phase | < 0.95 | Abbott IFU | Accredited
[6 empty] | Luteal Phase | 3.8 - 51
[6 empty] | Post Meno pausal | < 0.63
[6 empty] | 1st Trimester | 8.9 - 468
[6 empty] | 2nd Trimester | 72 - 303
[6 empty] | 3rd Trimester | 89 - 771
[6 empty] | Male | < 0.63 nmol/L
Prolactin | uIU/ml | Serum: Serum, Serum separator Plasma: Potassium EDTA, Sodium heparin, Lithium heparin | 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult Female | 109 - 557 | Abbott IFU | Accredited
[6 empty] | Adult Male | 73 -407
Prolactin, Macroprolactin | mIU/L | Serum: Serum, Serum separator Plasma: Potassium EDTA, Sodium heparin, Lithium heparin | 2-8°C: 7 days -20°C: 1 year | 24 hours | CALCULATION based on Abbott Alinity methodology for Prolactin and PEG precipitation | Adult Female | 79 - 347 | Abbott IFU | Accredited | FORMULA: (Prolactin post-PEG / Total Prolactin)*100 Macro-Prolactin is performed on Prolactin results above the normal range The post-PEG prolactin result indicates the approximate concentration of monomeric (biologically active) prolactin in the sample. Reference ranges are according to Beltran, L et al, Clinical Chemistry 54:10 1673-1681 (2008).
[6 empty] | Adult Male | 72 - 229
Rheumatoid Factor RF | IU/mL | Serum Serum tubes | 2-8°C: 2 days -20°C: 1 year | 24 hours | Abbott Alinity c, Immuno turbidimetric | Negative: | <30 | Abbott IFU | Accredited
Serum Protein Electrophoresis and Immunofixation |  | Serum : Serum tubes (with or without gel barrier) | 2-8°C: 10 days -20°C: 2 months | 10 days +2 days if Immunofixation required. | Electrophoresis: Sebia Capillarys 3 Octa capillary Immunofixation: Hydrasys 2 agarose gel |  | Fraction % Source: In-house study | Fraction g/L Source: In-house study | Accredited
[6 empty] | Albumin Fraction | 54.1 - 64.8 | 38 - 51
[6 empty] | Alpha - 1 Fraction | 3.1 - 5.2 | 2.3 - 3.9
[6 empty] | Alpha - 2 Fraction | 7.3 - 11.9 | 5.3 - 9.0
[6 empty] | Beta -1 Fraction | 5.1 - 7.8 | 3.6 - 5.7
[6 empty] | Beta- 2 Fraction | 3.6 - 7.7 | 2.6 - 5.9
[6 empty] | Gamma Fraction | 10.9 - 20.0 | 7.7 - 15.5
Sex Hormone Binding Globulin SHBG | nmol/L | Serum: Serum, Serum separator Plasma: Lithium heparin, Ammonium heparin, Sodium heparin | 2-8°C: 8 days -20°C: 3months | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult Female >19years | 20.0 - 155.0 | Abbott IFU | Accredited
[6 empty] | Pregnancy | <500
[6 empty] | Post Meno pausal | 26 - 118
[6 empty] | Adult Male >19years | 13.0 - 71.0
Sodium | mmol/L | Serum : Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium and Sodium heparin | 20-25°C: 2weeks 2-8°C: 2 weeks -20°C: 1 year | 24 hours | Abbott Alinity c, Indirect ISE | Adult | 136 - 145 | Abbott IFU | Accredited | Must be seperated ASAP
Sodium (Urine) | mmol/24hr | Urine (random; 24-hour) Without preservatives | 20-25°C: 45 day 2-8°C: 45 days -20°C: 1 year | 24 hours | Abbott Alinity c, Indirect ISE | Adult Male | 40 - 220 | Abbott IFU | Accredited
[6 empty] | Adult Female | 27 - 287
Tacrolimus | ug/L | Whole Blood: EDTA | 2-8°C: 7 days | 24 hours | Abbott Alinity I, Manual Pre-treatment precipitation. Chemiluminescent Microparticle Immunoassay (CMIA) | Target 24hr trough levels | 5 - 20 | Abbott IFU, Consensus document: therapeutic monitoring of tacrolimus (FK-206). Ther Drug Monit 995; 17(6): 606 - 14. | Accredited | The therapeutic range of tacrolimus is not clearly defined, but target 24-hour trough whole blood concentrations are 5 - 20 ug/L early post-transplant. Higher concentrations are associated with an increased incidence of adverse effects. 24-hour trough concentrations are 33 - 50% less than the corresponding 12-hour trough levels.
Testosterone | nmol/L | Serum: Serum, Serum separator Plasma: Dipotassium EDTA | 20-25°C: 8hrs 2-8°C: 7 days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Male 21 - 49 years | 8.33 - 30.19 | Abbott IFU | Accredited
[6 empty] | Male >50 years | 7.66 - 24.82
Free Androgen Index FAI | Ratio | Serum: Serum, Serum separator Plasma: Lithium heparin, | 2-8°C: 7 days -20°C: 3months | 24 hours | CALCULATION based on Abbott Alinity methodology for Testosterone and SHBG | Male: | 20.4 - 81.2 | Abbott IFU | Accredited | FORMULA: Free Androgen Index = (Testosterone*100) / SHBG
[6 empty] | Female (Pre-meno pausal): | 0.5 - 7.3
[6 empty] | Female (Post-meno pausal): | 0.6 - 8.0
Theophyline | mg/L | Serum: Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin Potassium EDTA Sodium citrate Sodium fluoride/potassium oxalate | 2-8°C: 3 month -20°C: 3 month | 24 hours | Abbott Alinity c, Enzyme Immunoassay | Adult | 8 - 20 | Abbott IFU | Accredited | Trough sample required.
Theophylline serum concentrations above 20 mg/L are often associated with toxicity.
TSH - Thyroid Stimulating Hormone | uIU/mL | Serum : Serum tubes Serum separator tubes . Plasma: Lithium heparin, Sodium heparin, Potassium EDTA | 2-8°C: 7 days -10°C: 6months | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Pediatric |  | Pediatric: CALIPER Adult: Abbott IFU | Accredited
[6 empty] | 4 day to < 6 months | 0.73 - 4.77
[6 empty] | 6 months to < 14 year | 0.70 - 4.17
[6 empty] | 14 to < 19 year | 0.47 - 3.41
Adult
[6 empty] | 19 - 120 year | 0.35 - 4.94
Thyroxine (TT4) | nmol/L | Serum: Serum, Serum separator Plasma: Potassium EDTA, Lithium heparin, Lithium heparin plasma separator, Sodium heparin | 2-8°C: 6 days -20°C: 6 days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 63 - 151 | Abbott IFU | Accredited
Total Prostate Specific Antigen PSA | µg/L | Serum : Serum, Serum separator | 2-8°C: 1 day -20°C: 24weeks | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult Male <50 years | < 2.0 | NCCP Prostate Cancer GP Referral Guideline v 5 2018; Abbott IFU | Accredited | Do not take sample within 1 week of digital rectal examination, or 6 weeks after prostate biopsy.
[6 empty] | Adult Male 50 - 59 years | < 3.0
[6 empty] | Adult Male 60 - 69 years | < 4.0
[6 empty] | Adult Male >70 years | < 5.0
Total Protein | g/L | Serum : Serum Serum separator Plasma Dipotassium EDTA Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c, Biuret Reaction | Adult, ambulatory | 64 - 83 | Abbott IFU | Accredited
Total T3 | nmol/L | Serum: Serum, Serum separator Plasma: Potassium EDTA, Lithium heparin, Sodium heparin | 2-8°C: 6 days -20°C: 6 days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 0.54 - 2.96 | Abbott IFU | Accredited
Transferrin | g/L | Serum: Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin EDTA | 20-25°C: 3 days 2-8°C: 3 days -20°C: 6 months | 24 hours | Abbott Alinity c, Immuno turbidimetric | Male 14 to 60 Years | 1.74 - 3.64 | Abbott IFU | Accredited
[6 empty] | Female 14 to 60 Years | 1.80 - 3.64
[6 empty] | Male 60 to 80 Years | 1.63 - 3.44
[6 empty] | Female 60 to 80 Years | 1.73 - 3.60
Transferrin Saturation | % | Serum : Serum tubes (with or without gel barrier) Plasma -Acceptable anticoagulants are: Sodium heparin Potassium EDTA Sodium citrate | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | CALCULATION based on Abbott Alinity methodologies for Iron and UIBC (Latent capacity) | Male Adult | 20.0 - 50.0 | Abbott IFU | Accredited | FORMULA: Transferrin Saturation = ((Iron*100) / (Iron + UIBC))
[6 empty] | Female Adult | 15.0 -50.0
Tricyclic Antidepressant (TCA) | ng/ml | Urine Clean plastic or glass container | 2-8°C: 2 days -20°C:longer | 24 hours | Nal Von Minden GmbH- Point of Care dipstick | Positivity Cut-off | 1000 | Nal Von Minden GmbH IFU | Unaccredited | The Drugs Of Abuse assays (with the exception of ethanol and urinary creatinine) are screening tests which provide preliminary positive or negative results. If clinically indicated, confirmation by GCMS is available on request. Specimen storage 2 days as per Kit insert: version 1.02 2023-05-22.
Triglyceride | mmol/L | Serum: Serum separator Plasma: Lithium heparin Lithium heparin separator Sodium heparin | 20-25°C: 2 days 2-8°C: 7 days -20°C: 3 months | 24 hours | Abbott Alinity c, Glycerinphosphate oxidase | Adult (ideal, fasting) | 0.6 - 1.7 | Abbott IFU, European Guidelines. http://www.eas-society.org/guidelines-2.aspx | Accredited | Fasting sample required.
Troponin-I, STAT high sensitive | pg/mL | Serum: Serum with and without separator Serum with thrombin-based clot activator Plasma: Lithium heparin with and without separator K2 EDTA K3 EDTA | 20-25°C: 8hrs 2-8°C: 24hours -20°C: 31 days | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Male | < 34 | Abbott IFU | Accredited
[6 empty] | Female | < 16
Total Iron Binding Capacity (TIBC) |  | Serum : Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | CALCULATION based on Abbott Alinity methodologies for Iron and UIBC (Latent capacity) | Adult | 44.8 - 76.1 | Abbott IFU | Accredited | Formula: TIBC = Iron + UIBC (LC)
Unsaturated Iron Binding Capacity UIBC | µmo/L | Serum : Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin | 20-25°C: 7 days 2-8°C: 3 weeks -20°C: 1 year | 24 hours | Abbott Alinity c, Ferene | Male | 12.4 - 43.0 | Abbott IFU | Accredited | Specimen should be collected in the morning to avoid low results due to diurnal variation. Drugs: Methyldopa and oxytetracycline cause artificially high UIBC values. Other: Pathologically high levels of albumin (7 g/l) decrease the apparent UIBC value significantly
[6 empty] | Female | 12.5 - 55.5
Urea | mmol/L | Serum : Serum tubes (with or without gel barrier) Plasma Collection tubes Acceptable anticoagulants are: Lithium heparin (with or without gel barrier) Sodium heparin | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c, Urease | Adult, Male <50years | 3.2 - 7.4 | Abbott IFU | Accredited
[6 empty] | Adult, Female <50years | 2.5 - 6.7
[6 empty] | Adult Male >50years | 3.0 - 9.2
[6 empty] | Adult Female >50years | 3.5 - 7.2
Urea - Urine | mmol/ 24hr | Urine (24 hour) Clean plastic or glass container with or without preservatives | 20-25°C: 7 days 2-8°C: 7 days -20°C: 1 year | 24 hours | Abbott Alinity c, Urease | Adult | 428 - 714 | Abbott IFU | Accredited
| mmol/L | Spot urine (random) |  | 24 hours |  | No reference range available
Uric acid | umol/L | Serum: Serum tubes Serum separator tubes Plasma Lithium heparin tubes Lithium heparin separator tubes Sodium heparin tubes | 20-25°C: 8 hour 2-8°C: 3 days -20°C: 3 months | 24 hours | Abbott Alinity c, Uricase | Male 13-79 years | 220 - 450 | Abbott IFU | Accredited
[6 empty] | Female 13-79 years | 150 - 370
Uric acid - Urine | umol/ 24hr | Urine (24 hour) Clean plastic or with or without preservatives. | 20-25°C: 2 days 2-8°C: 2 days | 24 hours | Abbott Alinity c, Uricase | Male (Purine Free Diet) | <2480 | Abbott IFU | Accredited
[6 empty] | Female (Purine Free Diet) | slighly lower
| umol / L | Urine (random specimens or timed specimens collected over intervals shorter than 24 hours) Clean plastic or glass container with or without preservatives. |  |  |  | Male (Low Purine Diet) | < 2830
[6 empty] | Female (Low Purine Diet) | < 2360
[6 empty] | High Purine Diet | < 5900
[6 empty] | Average Diet | 1480 - 4430
Urine Protein | mg/24hr | Urine (24hr/timed). Clean plastic or glass container without preservatives. | 20-25°C: 1 day 2-8°C: 7 days -20°C: 1 month | 24 hours | Abbott Alinity c, Benzethonium chloride | Adult | 0 - 300 | Abbott IFU | Accredited | Keep specimens on ice during collection. Testing of fresh urine specimens is suggested. Avoid collection of specimens within 24 hours of intense exercise since this can falsely elevate protein excretion.
Urinary Protein/ Creatinine Ratio |  |  |  | 24 hours | CALCULATION based on Abbott Alinity methodologies for urinary protein and urinary creatinine |  |  |  | Accredited | Formula: Protein: Creatinine = Urinary Protein / Urinary Creatinine
Urinary Ph (DOA) |  | Urine Clean plastic or glass container | 2-8°C: 5 days | 24 hours | Abbott Alinity c, DRI pH-Detect Test | Normal | 4.7 - 7.8 | DRI pH-Detect test IFU | Accredited
Urinary Creatinine (DOA) | mmol/L | Urine Clean plastic or glass container | 20-25°C: 2 days 2-8°C: 6 days -20°C: 6 months | 24 hours | Abbott Alinity c, Enzymatic | Normally concentrated urine | > 2.0 | EWDTS guidelines 2004; Abbott IFU | Accredited
[6 empty] | Dilute urine sample | 0.5 - 2.0
[6 empty] | Sample inegrity questionnable | < 0.5
Valproic acid | µg/mL | Serum : Serum tubes (with or without gel barrier) Plasma-Acceptable anticoagulants: Lithium heparin, Sodium heparin Potassium EDTA. Heparin gel plasmsa separator | 2-8°C: 2 days -20°C: 7 days | 24 hours | Abbott Alinity c, Particle-enhanced turbidimetric inhibition immunoassay (PETINIA) | Adult | 50 - 100 | Abbott IFU | Accredited
Vancomycin | µg/mL | Serum : Serum tubes (with or without gel barrier / clot activator) Plasma -Acceptable anticoagulants are: Sodium heparin, Lithium heparin K2-EDTA, K3-EDTA | 2-8°C: 7 days -10°C: 14 days | 24 hours | Abbott Alinity c, Homogeneous particle enhanced turbidimetric inhibition immunoassay (PETINIA) | Trough: | 5 - 20 | Abbott IFU, Therapeutic monitoring of vancomycin in adult patients Consensus review. Am J Health-Syst Pharm. 2009; 66:82 - 98.) | Accredited | A trough level of 15 - 20 ug/mL is recommended for certain infections: treatment of MRSA, hospital acquired pneumonia, bacterial meningitis & osteomyelitis. Vancomycin toxicity is not seen with a trough level of up to 20 ug/mL. Therapeutic peak serum levels of 20 to 40 μg/mL (13.80 to 27.60 μmol/L) and trough levels of 5 to 10 μg/mL (3.45 to 6.90 μmol/L) have been reported to be effective for most strains of staphylococci and streptococci.
[6 empty] | Peak: | 20 - 40
Vitamin B12 | pg/mL | Serum: Serum, Serum separator. Plasma: Lithium heparin plasma separator, Sodium heparin, Dipotassium EDTA | 20-25°C: 3 days 2-8°C: 7 days -20°C: longer | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 200 - 883 | Abbott IFU | Accredited | Fasting sample recommended. Please note that, with effect from the 01st of March 2025, due to uncertainty of measurement of vitamin B12 at lower levels, our reference range for Serum Vitamin B12 changed from 187 to 200.
Vitamin D, 25-OH | nmol/L | Serum: Serum, Serum separator Plasma: Dipotassium EDTA, Tripotassium EDTA Sodium heparin Lithium heparin powder Lithium heparin plasma separator | 20-25°C: 72hrs 2-8°C: 12 days -20°C: 1 year | 24 hours | Abbott Alinity i Chemiluminescent Microparticle Immunoassay (CMIA) | Adult | 30 - 125 | 1. Dietary reference intakes for calcium and Vitamin D. Washington, DC: The National Academies Press. 2. J Clin Endocrinol Metab, October 2011, 96(10):2987- 2996. 3. Abbott IFU | Accredited
25-OH-VITAMIN D CUTOFFS BASED ON DIETARY REFERENCE INTAKES
[6 empty] | Increased risk of deficiency | < 30
[6 empty] | Increased risk of in- adequacy | < 40
[6 empty] | Adequacy | > 50
[6 empty] | Increased risk of excess: | > 125